Shielding Calculation Report

Varian iX

Methodist Hospital ~ Willowbrook

August 22, 2007

By: Susan Richardson, Ph.D., D.A.B.R.

MP 10019

1. Introduction

This report reviews the radiation shielding for the therapy rooms in a new facility being built in Willowbrook, Texas, by the Methodist Hospital System.

The center will begin with one linear accelerator and a vault for a second possible accelerator. The vault that is being built out is the south vault and all calculation points will be relative to this vault.  
The calculations are based on the plans supplied by Carl Butler of Harrell Architects on 7/24/07. Any changes from the plans may invalidate the calculations.

2. Linear Accelerator

The accelerator is a Varian iX with On-Board-Imaging manufactured by Varian. The unit will be capable of producing 6 MV and 18 MV photons and electrons ranging from 6 MeV to 20 MeV. The accelerator is also capable of doing Intensity Modulated Radiation Therapy which will be discussed later. Electrons do not contribute to shielding requirements in the case of a linear accelerator that also produces photons.
The facility is attached to an office building, but there are no offices or useable space above the vaults.  Currently no buildings are near the vaults themselves other than the one shown in the elevation plan. Addition of new buildings near the vaults may invalidate calculations.  Floor and elevation plans used are found in Appendix 1.  

3. Shielding Design Goal
The shielding design goals (P) are 0.02 mSv/week (1mSv/yr) for uncontrolled areas and 0.1 mSv/week (5mSv/yr) for controlled areas.  

The maximum dose equivalent in any one hour (TADR) is 0.02 mSv (20µSv).

4. Assumptions

· The concrete density is 147 lbs./cu ft.
· Distances to calculation points are the distance from the target or isocenter to the wall + 1 foot. Ceiling distances are + 3 feet. Isocenter to target distance is 1 m.
· Primary and Leakage Tenth Value Layers (TVL) and Half Value Layers (HVL) are found below and are from NCRP Report #151 Tables B.2 and B.7.
	Energy (MV)
	Radiation Source
	TVL1
(first Tenth Value Layer)
	TVLe

(last Tenth Value Layer)
	HVL



	
	
	Cm
	Inches
	Cm
	Inches
	Cm
	Inches

	18
	Primary
	45
	17.7
	43
	16.9
	13.5
	5.3

	6
	Primary
	37
	14.6
	33
	13
	11.1
	4.37

	18
	Leakage
	36
	14.2
	34
	13.4
	10.8
	4.25

	6
	Leakage
	34
	13.4
	29
	11.4
	10.2
	4.02


· Workload is the dose (Gy) per week at Isocenter and it is assumed that the conventional workloads, Wconv  are W(6 MV) = 300 and W(18 MV) = 450  Gy/week.
IMRT considerations:

[image: image1.wmf]IMRT

conv

L

W

C

W

W

*

+

=


Assume 80% of 6 MV patients are IMRT, and 40% of 18 MV patients are IMRT. C is the monitor unit efficiency and is assumed to be 3.  Therefore the workloads with IMRT are 780 Gy/ week and 810 Gy/week respectively. Typically, conventional workloads will be used for primary barriers, and IMRT workloads will be used for scatter, leakage, and TADR calculations.
· Occupancy factors are given by table B1 of NCRP #151

5.  Summary (full calculations are given in Appendix 2)
	Point
	P
	inches  needed
	inches present
	ok?

	A
	Uncontrolled
	61.9 concrete
	116
	Y

	B
	Controlled
	67.2 concrete
	151
	Y

	C
	Uncontrolled
	25.9 concrete
	42
	Y

	D
	Uncontrolled
	20.67 concrete
	39
	Y

	E
	Controlled
	44.5 concrete
	44.5
	Y

	F
	Controlled
	32.5 concrete
	63.8
	Y

	G
	Uncontrolled
	27.4 concrete
	48
	Y

	H
	Uncontrolled
	67.5 concrete
	80
	Y

	I
	Uncontrolled
	28.2 concrete
	44
	Y

	DOOR
	Controlled
	3.1´´ BPE & 0.86´´ Pb
	0
	N



Locations of points are given below.
[image: image2.png]Area

LINAC | PointF
Control .

PointD &

o

artyard

Point A





6.  References
· National Council of Radiation Protection and Measurements. Structural Shielding Design and Evaluation for Megavoltage X- and Gamma-Ray Radiotherapy Facilities, Washington DC: NCRP, NCRP Report No. 151. 2006.  

· McGinley, PH. Shielding Techniques in Radiation Oncology Facilities, 2nd Edition. Medical Physics Publishing. Madison, WI. 2002.

· CLINAC 1800, 2100C(/D) & 2300C/D, 21EX, 23EX RADIATION LEAKAGE DATA (http://www.varian.com/onc/shared/pdf/12000.pdf)

· National Council of Radiation Protection and Measurements. Structural shielding design and evaluation for medical use of x-ray and gamma rays of energies up to 10 MeV. Washington DC: NCRP, NCRP Report No. 49; 1976.

· National Council of Radiation Protection and Measurements. Radiation protection design guidelines for 0.1-100 MeV particle accelerator facilities. Washington DC: NCRP, NCRP Report No. 51; 1977.

· 25 Texas Administrative Code §289.229. Radiation Safety Requirements for Accelerators, Therapeutic Radiation Machines, and Simulators. October, 2000.

Appendix 1

Building North Section:

[image: image3.png]JFie ot Vew Document Tods Advanced Windon Fep 5

Powen P [Blsme (e

©mEe - ®

- uthars sit

Bookmarks

To face of existing building

TopuuE
ey
P i d

T
AP

Signatures

ion_g

wawos o

ssowars_g

w-mer_g

[_commerts

20y 2007 9601 KATY FRIVY. 4270
METHODIST WILLOWBROOK HOSPITAL BUILDING SECTION HOUSTON, TEXAS 77024
CENTERFELD CTC North Section TEL 713722701
Houston, Toras FAX: 7137227072

Txiln__ < I

KK 101 b PO Ol





Floor Plan: Portrait

[image: image4.png]dobe Acrobat Professional

FLOOR PLAN- OPTION 2 Porrait b.pdf]
Fle Edt Vew Document Tools Advanced Window Help

D open @ [ seve

print (@, email ) search Create PDF ') Review & Comment ~

] Advanced Edting ~

®e% - ® | DY OF Hhowros

[ Paves Y Layers [ Signatures ff Bookmarks.

[_commerts

4]

ning /
Patiend|
18 oress.fl
Tollet fil outdoor
| Encioses
oo )
5 i
|l
et
e | CINAG
ialtng || Control ¢
1 oncoLoe|
i
55 (il ueAey
ageql 55 i Control Z|

Qutdoor 5
Enclosed |
[Caurtyard

FLOOR PLAN





Appendix II
· Point A
Description: Primary barrier along south wall

Occupancy factor (T):  Parking lot = 1/40 = 0.025

Distance from Isocenter to point, d= 23.5´+ 1´ = 24.5´= 7.46m

Dose per week outside barrier (P)= uncontrolled = 
[image: image5.wmf]6

10

*

20

-

Sv
Use Factor (U) = .25
[image: image6.png]29, e Ml
e
<7

L3 A

o

5% :

S

3
0 idtence: 2050
& PomtA @ 1
NE R 24 %
Lo 27

44"





· Primary Barrier Calculation

Transmission factor of the primary barrier (Bx):


[image: image7.wmf]4

2

6

2

10

*

09

.

5

25

.

*

025

.

*

450

)

1

46

.

7

(

*

10

*

20

)

18

(

-

-

=

+

=

=

WUT

Pd

MV

B

pri

x


Number of tenth value layers required (n):


[image: image8.wmf]29

.

3

10

*

09

.

5

1

log

1

log

)

log(

4

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

=

-

x

x

B

B

n


Barrier thickness (t) is given by:


[image: image9.wmf]"

5

.

56

6

.

143

43

*

)

1

29

.

3

(

45

)

1

(

1

=

=

-

+

=

-

+

=

cm

TVL

n

TVL

t

e


When a barrier is thickness t is greater than the first TVL, the total transmission factor, B, is given by:


[image: image10.wmf]

 EMBED Equation.3  [image: image11.wmf]ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

=

e

TVL

TVL

t

x

B

1

1

10


To determine if this is adequate for the additional workload from 6 MV, the following are used: 

[image: image12.wmf]week

Sv

week

Sv

d

WUT

B

MV

H

x

m

54

.

1

10

*

54

.

1

)

1

46

.

7

(

025

.

*

25

.

*

300

*

10

)

6

(

6

2

33

37

6

.

143

1

2

=

=

+

=

=

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-


The 6 MV dose equivalent is only 8% of the shielding design goal and will not effect the primary beam shielding.
· TADR Calculation
The maximum dose equivalent in any one hour is determined from the maximum absorbed dose rate at isocenter. This is 
[image: image13.wmf].

o

D

= 12 Gy/min or 720 Gy/hr.  The instantaneous dose rate at location A with transmission factor Bx is:


[image: image14.wmf](

)

hr

Sv

x

d

B

D

MV

IDR

x

3

2

4

2

0

.

10

*

12

.

5

1

46

.

7

10

09

.

5

*

720

*

)

18

(

-

-

=

+

=

=


And


[image: image15.wmf](

)

hr

Sv

MV

IDR

4

2

33

37

6

.

143

1

10

*

91

.

5

1

46

.

7

10

*

720

)

6

(

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

=

+

=


The weekly time averaged dose equivalent rate, Rw, is the time averaged dose rate at the location averaged over a 40 hour work week.  It is as follows:


[image: image16.wmf]week

Sv

week

Sv

D

WU

IDR

MV

R

w

m

800

10

*

00

.

8

720

25

.

*

450

*

10

*

12

.

5

*

)

18

(

4

3

.

0

=

=

=

=

-

-


And


[image: image17.wmf]week

Sv

week

Sv

MV

R

w

m

62

10

*

16

.

6

720

25

.

*

300

*

10

*

91

.

5

)

6

(

5

4

=

=

=

-

-


The maximum number of patients per hour is 6. The average number of patients per hour is 5.  M = 6/5=1.2.  The dose equivalent in any one hour from both 6MV and 18MV patients is:

[image: image18.wmf](

)

hour

Sv

R

M

R

total

x

h

m

26

62

800

*

40

2

.

1

40

)

(

=

+

÷

ø

ö

ç

è

æ

=

÷

ø

ö

ç

è

æ

=


This does not satisfy the 20µSv in any one hour requirement.  1 HVL (18MV) must be added to meet the requirement.  


[image: image19.wmf]"

9

.

61

1

.

157

5

.

13

6

.

143

=

=

+

=

cm

cm

cm

t


· Scatter Dose Calculation
For this patient scatter calculation, a position of 10º off the beam central axis is assumed. This is the worst case scenario because when the beam is pointing at point A the scatter fraction (α) is highest and the energy of the small angle scattered radiation is also the highest.  We will just consider 18 MV scatter due to the higher energy and penetration of the scatter beam.

Calculated slant thickness of barrier = 157.1 cm

Distance from target to patient, dsca = 1m

Distance from isocenter to point A = dsec= 7.46m

F = field area at mid depth of patient at 1m = 40x40 cm2
α = Scatter fraction for 10º= 1.42x10-2 from table B.4 of NCRP 151

TVL (18MV, scattered at 10º) = 45 cm concrete from table B.5a of NCRP 151

The maximum transmitted patient scattered dose equivalent at location A from 18 MV is:


[image: image20.wmf]4

45

1

.

157

10

*

23

.

3

10

10

-

÷

ø

ö

ç

è

æ

-

÷

ø

ö

ç

è

æ

-

=

=

=

TVL

t

sca

B


Then


[image: image21.wmf](

)

(

)

week

Sv

d

d

WUT

F

B

H

sca

sca

sca

/

10

*

27

.

9

46

.

7

*

1

025

.

*

25

.

*

450

400

40

*

40

 

10

*

1.42

10

*

23

.

3

400

7

2

2

2

-

4

2

sec

2

-

-

=

÷

ø

ö

ç

è

æ

÷

ø

ö

ç

è

æ

=

=

a


The dose of 0.927 µSv/week is well below the 20µSv/week limit and is therefore negligible. The primary barrier is sufficient for patient scattered radiation.
· Leakage Radiation Calculation
W = IMRT workloads given in part 4 - Assumptions
Use TVLs are the same as tabulated in part 4 – Assumptions
The transmitted 6 MV leakage dose equivalent is:


[image: image22.wmf]week

Sv

d

WT

B

H

L

L

L

/

10

*

1

.

1

46

.

7

)

025

)(.

780

(

10

10

10

*

*

3

2

3

29

34

1

.

157

1

2

3

m

-

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

=

=

=


And for 18MV:

[image: image23.wmf]week

Sv

d

WT

B

H

L

L

L

/

10

*

76

.

5

46

.

7

)

025

)(.

810

(

10

10

10

*

*

3

2

3

34

36

1

.

157

1

2

3

m

-

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

=

=

=


Both these values are well below the limit.  Because the treatment head is generally shielded to better than 0.1%, the primary barrier is adequate for shielding from additional leakage radiation, even with additional workload from IMRT. 
· Point B

Description: Primary barrier between vaults

Occupancy factor (T):   Adjacent Treatment Room = 0.5

Distance from Isocenter to point, d = 22.5´+ 1´= 23.5´ = 7.16m

Dose per week outside barrier (P) = controlled area = 
[image: image24.wmf]3

10

*

1

.

0

-

 Sv
Use Factor (U) = .25
[image: image25.png]



Transmission factor of the primary barrier (Bx):


[image: image26.wmf]4

2

3

2

10

*

18

.

1

25

.

*

5

.

*

450

)

1

16

.

7

(

*

10

*

1

.

0

)

18

(

-

-

=

+

=

=

WUT

Pd

MV

B

pri

x


Number of tenth value layers required (n):


[image: image27.wmf]9

.

3

10

*

18

.

1

1

log

1

log

4

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image28.wmf]"

2

.

67

8

.

170

43

*

)

1

9

.

3

(

45

)

1

(

1

=

=

-

+

=

-

+

=

cm

TVL

n

TVL

t

e


To determine if this is adequate for the additional workload from 6 MV, the following are used: 


[image: image29.wmf](

)

week

Sv

week

Sv

d

WUT

B

MV

H

x

m

95

.

4

10

*

95

.

4

1

16

.

7

5

.

*

25

.

*

300

*

10

)

6

(

6

2

33

37

8

.

170

1

2

=

=

+

=

=

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-


The 6 MV dose equivalent is only 5% of the shielding design goal and will not effect the primary beam shielding.

The instantaneous dose rate at point B with transmission factor Bx is:


[image: image30.wmf](

)

hr

Sv

x

d

B

D

MV

IDR

x

3

2

4

2

0

.

10

*

28

.

1

1

16

.

7

10

18

.

1

*

720

*

)

18

(

-

-

=

+

=

=



[image: image31.wmf](

)

hr

Sv

MV

IDR

5

2

33

37

8

.

170

1

10

*

51

.

9

1

16

.

7

10

*

720

)

6

(

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

=

+

=


Rw is as follows:


[image: image32.wmf]week

Sv

week

Sv

D

WU

IDR

MV

R

w

m

200

10

*

00

.

2

720

25

.

*

450

*

10

*

28

.

1

*

)

18

(

4

3

.

0

=

=

=

=

-

-



[image: image33.wmf]week

Sv

week

Sv

MV

R

w

m

9

.

9

10

*

9

.

9

720

25

.

*

300

*

10

*

51

.

9

)

6

(

6

5

=

=

=

-

-


The dose equivalent in any one hour from both 6MV and 18MV patients is:


[image: image34.wmf](

)

hour

Sv

R

M

R

total

x

h

m

3

.

6

9

.

9

200

*

40

2

.

1

40

)

(

=

+

÷

ø

ö

ç

è

æ

=

÷

ø

ö

ç

è

æ

=


This satisfies the 20µSv in any one hour requirement.  

· Patient Scatter and Leakage Calculation
As we have seen from point A, patient scatter and leakage doses are negligible when the point in question is behind a primary barrier. Therefore, the patient scatter and leakage calculations will not be duplicated for point B.

· Point C

Description: Secondary barrier along East wall directly behind Linac

Occupancy Factor:  Parking Lot = 1/40 =.025

Use Factor = 1 for scatter barriers

W = IMRT workloads given in part 4 - Assumptions 
Distance from target to patient, dsca = 1m

Distance from isocenter to point C = dsec= dL=16´+1´=17´=5.18m
Dose per week outside barrier (P)= uncontrolled = 
[image: image35.wmf]6

10

*

20

-

 Sv
[image: image36.png]



·  Patient Scatter Calculation
α (18 MV)= Scatter fraction for 90º= 1.89x10-4 from table B.4 of NCRP 151

α (6 MV)= Scatter fraction for 90º= 4.26 x10-4 from table B.4 of NCRP 151

TVL (6 MV, scattered at 90º) = 17 cm concrete from table B.5a of NCRP 151

TVL (18 MV, scattered at 90º) = 19 cm concrete from table B.5a of NCRP 151

F = field area at mid depth of patient at 1m = 40x40 cm2
For 6 MV the barrier thickness required is:

[image: image37.wmf](

)

(

)

(

)

(

)

(

)

2

4

-

2

2

6

2

2

sec

10

*

62

.

1

40

*

40

400

025

.

1

780

10

*

4.26

18

.

5

1

10

*

20

400

)

6

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:

[image: image38.wmf]79

.

1

10

*

62

.

1

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image39.wmf]cm

TVL

n

t

4

.

30

17

*

79

.

1

*

=

=

=


For 18 MV x-rays we get:


[image: image40.wmf](

)

(

)

(

)

(

)

(

)

2

4

-

2

2

6

2

2

sec

10

*

51

.

3

40

*

40

400

025

.

1

810

10

*

1.89

18

.

5

1

10

*

20

400

)

18

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:


[image: image41.wmf]46

.

1

10

*

51

.

3

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image42.wmf]cm

TVL

n

t

7

.

27

19

*

46

.

1

*

=

=

=


Since the scatter thicknesses required for 6 MV and 18MV are within 1 TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover patient scattered dose is:


[image: image43.wmf]cm

TVL

t

t

1

.

36

)

19

)(

301

(.

4

.

30

)

(

301

.

0

'

=

+

=

+

=


· Leakage Calculation

[image: image44.wmf](

)

2

3

2

6

3

2

10

*

75

.

2

025

.

*

780

*

10

18

.

5

10

*

20

*

10

*

)

6

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image45.wmf]56

.

1

10

*

75

.

2

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image46.wmf]cm

t

3

.

50

29

*

)

1

56

.

1

(

34

=

-

+

=


And for 18MV:


[image: image47.wmf](

)

2

3

2

6

3

2

10

*

65

.

2

025

.

*

810

*

10

18

.

5

10

*

20

*

10

*

)

18

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image48.wmf]58

.

1

10

*

65

.

2

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image49.wmf]cm

t

6

.

55

34

*

)

1

58

.

1

(

36

=

-

+

=


Since the leakage thicknesses required for 6 MV and 18MV are within one TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover leakage dose is:


[image: image50.wmf]"

9

.

25

8

.

65

)

34

)(

301

(.

6

.

55

)

(

301

.

0

'

=

=

+

=

+

=

cm

TVL

t

t


The scattered radiation contributions are negligible and this wall can be shielded for IMRT leakage alone.   
· TADR Calculation
The instantaneous dose equivalent rates for patient scatter are:

[image: image51.wmf]hr

Sv

d

F

D

MV

IDR

TVL

t

sca

PS

6

17

8

.

65

2

4

2

0

.

10

*

12

.

6

10

*

18

.

5

*

400

40

*

40

*

10

*

26

.

4

*

720

10

*

400

)

6

(

-

÷

ø

ö

ç

è

æ

-

-

÷

ø

ö

ç

è

æ

-

=

=

=

a

 and
[image: image52.wmf]hr

Sv

d

F

D

MV

IDR

TVL

t

sca

PS

6

19

8

.

65

2

4

2

0

.

10

*

95

.

6

10

*

18

.

5

*

400

40

*

40

*

10

*

89

.

1

*

720

10

*

400

)

18

(

-

÷

ø

ö

ç

è

æ

-

-

÷

ø

ö

ç

è

æ

-

=

=

=

a


The instantaneous dose equivalent rates for leakage we have:

[image: image53.wmf]hr

Sv

MV

IDR

Leak

4

2

29

34

8

.

65

1

3

10

*

14

.

2

18

.

5

10

*

10

*

720

)

6

(

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

=

=


and


[image: image54.wmf]hr

Sv

MV

IDR

Leak

4

2

34

36

8

.

65

1

3

10

*

56

.

3

18

.

5

10

*

10

*

720

)

18

(

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

=

=


We find the scattered radiation contributions are insignificant at point C, so only leakage radiation contributions will be computed.  The results are:


[image: image55.wmf]week

Sv

week

Sv

D

W

IDR

MV

R

Leak

w

m

232

10

*

32

.

2

720

780

*

10

*

14

.

2

*

)

6

(

4

4

.

0

=

=

=

=

-

-


And


[image: image56.wmf]week

Sv

week

Sv

D

W

IDR

MV

R

Leak

w

m

400

10

*

00

.

4

720

810

*

10

*

56

.

3

*

)

18

(

4

4

.

0

=

=

=

=

-

-


And therefore:

[image: image57.wmf](

)

hour

Sv

R

M

R

total

x

h

m

19

232

400

*

40

2

.

1

40

)

(

=

+

÷

ø

ö

ç

è

æ

=

÷

ø

ö

ç

è

æ

=


This is below the 20µSv/hour limit.

· Point D

Description: Outdoor courtyard secondary barrier along west wall 

Occupancy Factor:  Outside, no seating = 1/40 =.025

Use Factor = 1 for scatter barriers

W = IMRT workloads given in part 4 - Assumptions 
Distance from target to patient, dsca = 1m

Distance from isocenter to point D = dsec= dL=25.67´+1´=26.67´=8.13m

Dose per week outside barrier (P)= uncontrolled = 
[image: image58.wmf]6

10

*

20

-

 Sv
[image: image59.png]Point D #|
g

Qutdoor
Enclosed
Courtyard





As seen for point C, patient scatter will be negligible at this point. We will only consider leakage dose:

· Leakage Calculation

[image: image60.wmf](

)

2

3

2

6

3

2

10

*

78

.

6

025

.

*

780

*

10

13

.

8

10

*

20

*

10

*

)

6

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image61.wmf]17

.

1

10

*

78

.

6

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image62.wmf]cm

t

9

.

38

29

*

)

1

17

.

1

(

34

=

-

+

=


And for 18MV:


[image: image63.wmf](

)

2

3

2

6

3

2

10

*

53

.

6

025

.

*

810

*

10

13

.

8

10

*

20

*

10

*

)

18

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image64.wmf]18

.

1

10

*

53

.

6

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image65.wmf]cm

t

3

.

42

34

*

)

1

18

.

1

(

36

=

-

+

=


Since the leakage thicknesses required for 6 MV and 18MV are within one TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover leakage dose is:


[image: image66.wmf]"

7

.

20

5

.

52

)

34

)(

301

(.

3

.

42

)

(

301

.

0

'

=

=

+

=

+

=

cm

TVL

t

t


Similar to point C, the instantaneous dose rate is below the limit.

· Point E

Description: Directly outside treatment room door 

Occupancy Factor: control room door = 1/8=.125
Distance from isocenter to point F=dL=20.39´+1´=21.39´=6.2m
Dose per week outside barrier (P)= controlled = 
[image: image67.wmf]3

10

*

1

.

0

-

 Sv
[image: image68.png][Distance: 203

EEED 7

KOULLEY|
H3HONED





· Leakage Calculation
As seen for previous points, patient scatter will be negligible at this point. We will only consider leakage dose.  Point E is complicated by the fact that dose through the door is also at this point and contributes to the 0.1mSv dose limit.  Instead of calculating required barrier thickness, we will calculate the actual expected dose through the planned barrier. The existing barrier slant thickness, t, is 113.1 cm (44.5´´).

[image: image69.wmf]week

Sv

d

WT

d

BWT

MV

H

e

TVL

TVL

t

m

47

.

0

2

.

6

125

.

*

780

*

10

*

10

10

10

10

)

6

(

2

29

34

1

.

113

1

3

2

1

3

2

3

1

=

=

=

=

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

-

-


And


[image: image70.wmf]week

Sv

d

WT

d

BWT

MV

H

e

TVL

TVL

t

m

42

.

1

2

.

6

125

.

*

810

*

10

*

10

10

10

10

)

18

(

2

34

36

1

.

113

1

3

2

1

3

2

3

1

=

=

=

=

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

-

ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

-

-


Adding together, we get 1.85µSv/week.  We will find in the dose calculation for the door that compensating for the additional shielding by the steel encasing of the door will net us an extra 16 µSv/week that can be allotted for this dose requirement.
· Point F

Description: Linac control room 

Occupancy Factor:  control room = 1

Distance from isocenter to point F=dL=26.26´+1´=27.26´=8.31m

Dose per week outside barrier (P)= controlled = 
[image: image71.wmf]3

10

*

1

.

0

-

 Sv
As seen for previous points, patient scatter will be negligible at this point. We will only consider leakage dose:

[image: image72.png]—

LINAC
Contral





· Leakage Calculation

[image: image73.wmf](

)

3

3

2

3

3

2

10

*

85

.

8

1

*

780

*

10

31

.

8

10

*

1

.

*

10

*

)

6

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image74.wmf]05

.

2

10

*

85

.

8

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image75.wmf]cm

t

5

.

64

29

*

)

1

05

.

2

(

34

=

-

+

=


And for 18MV:


[image: image76.wmf](

)

3

3

2

3

3

2

10

*

53

.

8

1

*

810

*

10

31

.

8

10

*

1

.

*

10

*

)

18

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image77.wmf]07

.

2

10

*

53

.

8

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image78.wmf]cm

t

4

.

72

34

*

)

1

07

.

2

(

36

=

-

+

=


Since the leakage thicknesses required for 6 MV and 18MV are within one TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover leakage dose is:


[image: image79.wmf]"

5

.

32

6

.

82

)

34

)(

301

(.

4

.

72

)

(

301

.

0

'

=

=

+

=

+

=

cm

TVL

t

t


Similar to point C, the instantaneous dose rate is below the limit.

· Point G

Description: Off primary wall at 30º away from primary beam
Occupancy Factor:  Parking Lot = 0.025

Use Factor = ¼ = .25

Distance from isocenter to point G =dsec= dL =18.26´+1´=19.26´=5.87m

Distance from target to isocenter = dsca = 1m

Dose per week outside barrier (P)= uncontrolled = 
[image: image80.wmf]6

10

*

20

-

 Sv
α (18 MV)= Scatter fraction for 30º= 2.53x10-3 from table B.4 of NCRP 151

α (6 MV)= Scatter fraction for 30º= 2.77 x10-3 from table B.4 of NCRP 151

TVL (6 MV, scattered at 30º) = 26 cm concrete from table B.5a of NCRP 151

TVL (18 MV, scattered at 30º) = 32 cm concrete from table B.5a of NCRP 151

F = field area at mid depth of patient at 1m = 40x40 cm2
[image: image81.png][Distance: 18.26

T





· Patient Scatter Calculation
For 6 MV the barrier thickness required is:


[image: image82.wmf](

)

(

)

(

)

(

)

(

)

2

3

-

2

2

6

2

2

sec

10

*

28

.

1

40

*

40

400

025

.

25

.

780

10

*

2.77

87

.

5

1

10

*

20

400

)

6

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:


[image: image83.wmf]89

.

1

10

*

28

.

1

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image84.wmf]cm

TVL

n

t

1

.

49

26

*

89

.

1

*

=

=

=


For 18 MV x-rays we get:


[image: image85.wmf](

)

(

)

(

)

(

)

(

)

2

3

-

2

2

6

2

2

sec

10

*

35

.

1

40

*

40

400

025

.

25

.

810

10

*

2.53

87

.

5

1

10

*

20

400

)

18

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:


[image: image86.wmf]87

.

1

10

*

35

.

1

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image87.wmf]cm

TVL

n

t

8

.

59

32

*

87

.

1

*

=

=

=


Since the scatter thicknesses required for 6 MV and 18MV are within 1 TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover patient scattered dose is:


[image: image88.wmf]"

4

.

27

5

.

69

)

32

)(

301

(.

8

.

59

)

(

301

.

0

'

=

=

+

=

+

=

cm

TVL

t

t


· Leakage Calculation

[image: image89.wmf](

)

2

3

2

6

3

2

10

*

53

.

3

025

.

*

780

*

10

87

.

5

10

*

20

*

10

*

)

6

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image90.wmf]45

.

1

10

*

53

.

3

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image91.wmf]cm

t

1

.

47

29

*

)

1

45

.

1

(

34

=

-

+

=


And for 18MV:


[image: image92.wmf](

)

2

3

2

6

3

2

10

*

40

.

3

025

.

*

810

*

10

87

.

5

10

*

20

*

10

*

)

18

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image93.wmf]47

.

1

10

*

40

.

3

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image94.wmf]cm

t

9

.

51

34

*

)

1

47

.

1

(

36

=

-

+

=


Since the leakage thicknesses required for 6 MV and 18MV are within one TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover leakage dose is:


[image: image95.wmf]cm

TVL

t

t

2

.

62

)

34

)(

301

(.

9

.

51

)

(

301

.

0

'

=

+

=

+

=


Similar to point C, the instantaneous dose rate is below the limit.

· Point H

Description: Primary ceiling barrier 

Occupancy factor (T):  roof = 1/40 = 0.025

Distance from Isocenter to point, d = 14.57´ + 1´ = 15.57´= 4.75m

Dose per week outside barrier (P) = uncontrolled = 
[image: image96.wmf]6

10

*

20

-

 Sv
Use Factor (U) = .25

The roof is unoccupied open space and the primary beam currently does not point towards another building.  The roof area is accessible by ladder only.  It will occasionally have maintenance personnel working on the roof. If the amount of time spent on the roof is more than approximately 2 hrs/month during operational hours, the radiation oncology physicist should be notified.  Furthermore, the barrier must be 11´ wide and run the full length of the roof in the direction of the primary beam swath.  
[image: image97.png]



· Primary Barrier Calculation

Transmission factor of the primary barrier (Bx):


[image: image98.wmf]4

2

6

2

10

*

35

.

2

25

.

*

025

.

*

450

)

1

75

.

4

(

*

10

*

20

)

18

(

-

-

=

+

=

=

WUT

Pd

MV

B

pri

x


Number of tenth value layers required (n):


[image: image99.wmf]63

.

3

10

*

35

.

2

1

log

1

log

)

log(

4

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

=

-

x

x

B

B

n


Barrier thickness (t) is given by:


[image: image100.wmf]"

2

.

62

158

43

*

)

1

63

.

3

(

45

)

1

(

1

=

=

-

+

=

-

+

=

cm

TVL

n

TVL

t

e


When a barrier is thickness t is greater than the first TVL, the total transmission factor, B, is given by:


[image: image101.wmf]

 EMBED Equation.3  [image: image102.wmf]ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

=

e

TVL

TVL

t

x

B

1

1

10


To determine if this is adequate for the additional workload from 6 MV, the following are used: 


[image: image103.wmf](

)

week

Sv

week

Sv

d

WUT

B

MV

H

x

m

21

.

1

10

*

21

.

1

1

75

.

4

025

.

*

25

.

*

300

*

10

)

6

(

6

2

33

37

158

1

2

=

=

+

=

=

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-


The 6 MV dose equivalent is only 6% of the shielding design goal and will not effect the primary beam shielding.

· TADR Calculation
The maximum dose equivalent in any one hour is determined from the maximum absorbed dose rate at isocenter. This is 
[image: image104.wmf].

o

D

= 12 Gy/min or 720 Gy/hr.  The instantaneous dose rate at location A with transmission factor Bx is:


[image: image105.wmf](

)

hr

Sv

d

B

D

MV

IDR

x

3

2

4

2

0

.

10

*

12

.

5

1

75

.

4

10

*

35

.

2

*

720

*

)

18

(

-

-

=

+

=

=


And


[image: image106.wmf](

)

hr

Sv

MV

IDR

4

2

33

37

158

1

10

*

68

.

4

1

75

.

4

10

*

720

)

6

(

-

þ

ý

ü

î

í

ì

÷

ø

ö

ç

è

æ

-

+

-

=

+

=


The weekly time averaged dose equivalent rate, Rw, is the time averaged dose rate at the location averaged over a 40 hour work week.  It is as follows:


[image: image107.wmf]week

Sv

week

Sv

D

WU

IDR

MV

R

w

m

800

10

*

00

.

8

720

25

.

*

450

*

10

*

12

.

5

*

)

18

(

4

3

.

0

=

=

=

=

-

-



[image: image108.wmf]week

Sv

week

Sv

MV

R

w

m

49

10

*

88

.

4

720

25

.

*

300

*

10

*

68

.

4

)

6

(

5

4

=

=

=

-

-


The maximum number of patients per hour is 6. The average number of patients per hour is 5.  M = 6/5=1.2.  The dose equivalent in any one hour from both 6MV and 18MV patients is:


[image: image109.wmf](

)

hour

Sv

R

M

R

total

x

h

m

5

.

25

49

800

*

40

2

.

1

40

)

(

=

+

÷

ø

ö

ç

è

æ

=

÷

ø

ö

ç

è

æ

=


This does not satisfy the 20µSv in any one hour requirement.  1 HVL must be added to meet the requirement.  


[image: image110.wmf]"

5

.

67

5

.

171

5

.

13

158

=

=

+

=

cm

cm

cm

t


·    Point I
Description: Secondary ceiling barrier 

Occupancy factor (T):  roof = 1/40 = 0.025

Distance from Isocenter to point (dsec): 16.74´ + 1´ = 17.74´ = 5.4m
Dose per week outside barrier (P)= uncontrolled = 
[image: image111.wmf]6

10

*

20

-

 Sv
Use Factor (U) = .25

α (18 MV)= Scatter fraction for 30º= 2.53x10-3 from table B.4 of NCRP 151

α (6 MV)= Scatter fraction for 30º= 2.77 x10-3 from table B.4 of NCRP 151

TVL (6 MV, scattered at 30º) = 26 cm concrete from table B.5a of NCRP 151

TVL (18 MV, scattered at 30º) = 32 cm concrete from table B.5a of NCRP 151

F = field area at mid depth of patient at 1m = 40x40 cm2
IMRT workloads are used.

[image: image112.png]



· Patient Scatter Calculation
For 6 MV the barrier thickness required is:


[image: image113.wmf](

)

(

)

(

)

(

)

(

)

2

3

-

2

2

6

2

2

sec

10

*

08

.

1

40

*

40

400

025

.

25

.

780

10

*

2.77

4

.

5

1

10

*

20

400

)

6

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:


[image: image114.wmf]97

.

1

10

*

08

.

1

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image115.wmf]cm

TVL

n

t

2

.

51

26

*

97

.

1

*

=

=

=


For 18 MV x-rays we get:


[image: image116.wmf](

)

(

)

(

)

(

)

(

)

2

3

-

2

2

6

2

2

sec

10

*

14

.

1

40

*

40

400

025

.

25

.

810

10

*

2.53

4

.

5

1

10

*

20

400

)

18

(

-

-

=

÷

ø

ö

ç

è

æ

=

=

F

WUT

d

Pd

MV

B

sca

ps

a


Then the number of TVLs is:


[image: image117.wmf]94

.

1

10

*

14

.

1

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image118.wmf]cm

TVL

n

t

08

.

62

32

*

94

.

1

*

=

=

=


Since the scatter thicknesses required for 6 MV and 18MV are within 1 TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover patient scattered dose is:


[image: image119.wmf]"

2

.

28

7

.

71

)

32

)(

301

(.

08

.

62

)

(

301

.

0

'

=

=

+

=

+

=

cm

TVL

t

t


· Leakage Calculation

[image: image120.wmf](

)

2

3

2

6

3

2

10

*

99

.

2

025

.

*

780

*

10

4

.

5

10

*

20

*

10

*

)

6

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image121.wmf]52

.

1

10

*

99

.

2

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image122.wmf]cm

t

2

.

49

29

*

)

1

52

.

1

(

34

=

-

+

=


And for 18MV:


[image: image123.wmf](

)

2

3

2

6

3

2

10

*

88

.

2

025

.

*

810

*

10

4

.

5

10

*

20

*

10

*

)

18

(

-

-

-

-

=

=

=

WT

d

P

MV

B

L

L


Then the number of TVLs is:


[image: image124.wmf]54

.

1

10

*

88

.

2

1

log

1

log

2

=

÷

ø

ö

ç

è

æ

=

÷

÷

ø

ö

ç

ç

è

æ

=

-

x

B

n


Barrier thickness (t) is given by:


[image: image125.wmf]cm

t

4

.

54

34

*

)

1

54

.

1

(

36

=

-

+

=


Since the leakage thicknesses required for 6 MV and 18MV are within one TVL of each other, one HVL needs to be added according to the two source rule.  Therefore, the total shielding to cover leakage dose is:


[image: image126.wmf]cm

TVL

t

t

6

.

64

)

34

)(

301

(.

4

.

54

)

(

301

.

0

'

=

+

=

+

=


Similar to point C, the instantaneous dose rate is below the limit.

· Maze Door

The dose at the treatment vault door is due to a contribution of photons and neutron doses. The dose equivalent at the door, Hd is given by:

[image: image127.wmf]LT

PS

LS

s

d

H

H

H

fH

H

+

+

+

=


Each contribution will be calculated separately and then summed together.

· Wall Scattered Radiation Component, HS.
The ceiling is 15´10´´ = 4.83m

Normal use workloads will be used for this calculation.

U = 0.25
dG = perpendicular distance from the target to the first reflection surface = 15.58´ = 4.75 m
dr = distance from the beam center at the first reflection surface, past the edge of the inner maze wall, to the midline of the maze =19.55´=5.96m

dz = centerline distance along the maze to the door =17.35´=5.29m

A0 = maximum field size projected onto scatter wall = 40x40 projected a distance of 4.75 m = 1.9 x 1.9 m2 = 3.8 m2.
Az = cross sectional area of the maze inner entry projected from A0 = 7.88´x 15´10´´ = 2.4m x 4.83m =11.6 m2. 

αG(18MV)= 1.6x10-3 from Table B.8a of NCRP 151, normal incidence, 75º angle of reflection.

αG(6MV)= 2.7x10-3 from Table B.8a of NCRP 151, normal incidence, 75º angle of reflection.

αz= 8x10-3 from Table B.8a of NCRP 151, normal incidence, 75º angle of reflection of 0.5 MeV.

[image: image128.png]E o SV
B %
g 4@
4 . 5
H
& = ;
 m— — - iz .
8 ll =
H "
af= :
RoRT R
iy g e :





Then HS is:


[image: image129.wmf](

)

(

)

(

)

[

]

(

)

week

Sv

d

d

d

A

A

WU

H

z

r

G

Z

Z

S

m

a

a

01

.

6

29

.

5

*

96

.

5

*

75

.

4

6

.

11

*

10

*

8

*

8

.

3

*

25

.

*

10

*

7

.

2

*

300

10

*

6

.

1

*

450

2

3

3

3

2

0

0

=

+

=

=

-

-

-


· Single Scattered Head Leakage Component, HLS.
Head leakage can scatter once against the wall and reach the maze door.  IMRT workloads will be used for this calculation. 

U=.25

Lf = head leakage radiation ratio =  0.1% per IEC 2002 requirement

A1 = area of wall seen from maze door = 11.33´ x 15.83´ = 16.63 m2
dzz = centerline distance along maze = 9.26m
dLS = distance from target to wall along maze centerline = 8.59 m
α1 (6MV) = reflection coefficient for scatter of leakage radiation from wall (from Table B.8b from NCRP 151) for 45º incidence and 0 º reflection = 6.4x10-3. 

α1 (18MV) = reflection coefficient for scatter of leakage radiation from wall (from Table B.8b from NCRP 151) for 45º incidence and 0 º reflection = 4.5x10-3. 

[image: image130.png]L]
f
|
|
|
|
: U e
s i -
3 ]
: |
3y 3
Lo B
= S
s,

Az

[Distance: 30.37 ft|





The dose equivalent at the maze door from head leakage is thus:


[image: image131.wmf](

)

(

)

(

)

[

]

(

)

week

Sv

d

d

A

WU

L

H

zz

f

LS

m

a

67

.

5

26

.

9

*

59

.

8

63

.

16

*

25

.

*

10

*

5

.

4

*

810

10

*

4

.

6

*

780

*

10

2

3

3

3

2

sec

1

1

=

+

=

=

-

-

-


· Patient Scattered Radiation Component, HPS.
Normal workloads will be used for this calculation.

U = .25

F = 40x40 cm2
A1 = area of wall seen from maze door = 11.33´ x 15.83´ = 16.63 m2
dzz = centerline distance along maze = 9.26m

dsca = 1 m

dsec = distance from isocenter to wall along maze centerline = 26.57´=8.1m

a1(6MV) = The 6 MV scatter fraction at 45º scatter angle (from Table B.4 from NCRP 151) = 1.39x10-3. 

a1(18MV) = The 18 MV scatter fraction at 45º scatter angle (from Table B.4 from NCRP 151) = 8.64x10-4. 

α1 = the concrete wall reflection coefficient for incident angle 45º and reflection angle of 0º for 0.5 MeV monoenergetic photons (NCRP 151 table B.8b)  = 2.22x10-2
[image: image132.png]ONCHRETE, + "
< BARTITION,

dax

[Distance: 30.37 ft|

g





The 6 and 18 MV patient scatter contributions are: 


[image: image133.wmf](

)

(

)

(

)

[

]

(

)

=

+

=

=

-

-

-

2

2

4

3

2

sec

1

1

1

.

8

*

26

.

9

*

1

63

.

16

*

10

*

22

.

2

*

400

40

*

40

*

25

.

*

10

*

64

.

8

*

810

10

*

39

.

1

*

780

400

)

(

zz

sca

PS

d

d

d

A

F

WU

a

H

a

q

 
[image: image134.wmf]week

Sv

H

PS

m

116

=


· Head Leakage Radiation Through the Maze Wall, HLT.
IMRT workloads will be used for this calculation.

U = .25

Lf = head leakage radiation ratio =  0.1% per IEC 2002 requirement

dL = distance from target to center of maze door through the inner maze wall = 21.66´=6.6m
t = the oblique wall slant thickness = 2.89´= 88 cm
TVLs are as given in Assumptions.

[image: image135.png]Distance: 21.66

Ao NOLIWY
+°313H0N0D .,





Calculating 18MV first, we find the barrier transmission for slant thickness t is:

[image: image136.wmf]3

34

36

88

1

1

10

*

96

.

2

10

10

)

18

(

1

-

þ

ý

ü

î

í

ì

-

+

-

ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

=

=

=

e

TVL

TVL

t

x

MV

B


Then the dose equivalent is

[image: image137.wmf]week

Sv

d

WUB

L

MV

H

L

f

LT

m

7

.

13

6

.

6

10

*

96

.

2

*

25

.

*

810

*

10

)

18

(

2

3

3

2

=

=

=

-

-


Similarly for 6 MV:


[image: image138.wmf]3

29

34

88

1

1

10

*

37

.

1

10

10

)

6

(

1

-

þ

ý

ü

î

í

ì

-

+

-

ï

þ

ï

ý

ü

ï

î

ï

í

ì

÷

÷

ø

ö

ç

ç

è

æ

-

+

-

=

=

=

e

TVL

TVL

t

x

MV

B


Then the dose equivalent is


[image: image139.wmf]week

Sv

d

WUB

L

MV

H

L

f

LT

m

15

.

6

6

.

6

10

*

37

.

1

*

25

.

*

780

*

10

)

6

(

2

3

3

2

=

=

=

-

-


Total dose equivalent at the wall due to leakage through the inner maze wall is 13.7 + 6.15 = 20µSv/week.
The fraction of the primary radiation transmitted through the patient is represented by f and is assumed to be 0.34 for a high energy Linac.  When use factors are employed for dose equivalent considerations, the total is not 4*Htot. It is estimated as 2.64*Htot (McGinley, 2002).The total dose equivalent due to scatter and leakage radiations, Htot is then:


[image: image140.wmf](

)

(

)

week

Sv

H

H

H

fH

H

H

LT

PS

LS

s

d

tot

m

389

20

116

67

.

5

01

.

6

*

64

.

2

64

.

2

*

64

.

2

=

+

+

+

=

+

+

+

=

=


· Neutron Capture Gamma-Ray Dose Equivalent at Maze Door
Photoneutron production and neutron capture is proportional to the leakage radiation workload of high-energy x-rays (18 MV).  

First, the area of the room must be known.

Average ceiling height = 14´= 4.27m
Average room length = 23.8´= 7.25m
Average room width = 25.5´ = 7.77m 

[image: image141.wmf](

)

(

)

(

)

{

}

2

241

25

.

7

*

77

.

7

77

.

7

*

27

.

4

25

.

7

*

27

.

4

*

2

m

S

r

=

+

+

=


[image: image142.png]Distance: 21.57





· Total neutron fluence per isocenter x-ray Gray at inner maze point, ΦA:
Qn = neutron source strength in neutrons emitted from the accelerator head per gray of x-ray absorbed dose at the isocenter = 1.0*1012 for a Varian iX model.
d1 = distance from isocenter to center of maze = 6.57m

β = neutron transmission through linac head = 1 for Lead housing


[image: image143.wmf]2

9

12

12

2

12

2

1

10

*

27

.

6

241

*

2

10

*

1.0

*

3

.

1

241

*

2

10

*

1.0

*

4

.

5

57

.

6

*

4

 

10

*

1.0

2

3

.

1

2

4

.

5

4

m

neutrons

S

Q

S

Q

d

Q

r

n

r

n

n

A

=

+

+

=

+

+

=

F

p

p

p

p

p

b

p

b


· Photon Dose Equivalent Calculation at the Maze Door, hφ:
K = ration of neutron capture gamma-ray dose equivalent to the total neutron fluence  = 6.9*10-16 Sv/m2 (NCRP 151)

ΦA= neutron fluence, found above

d2=distance along maze = 21.7´= 6.61m

TVD = Tenth Value Distance for photons = 5.4 m for x-rays from 18-25 MV. 


[image: image144.wmf]Gy

Sv

K

h

TVD

d

A

7

4

.

5

61

.

6

9

16

10

*

58

.

2

10

*

10

*

27

.

6

*

10

*

9

.

6

10

2

-

÷

ø

ö

ç

è

æ

-

-

÷

ø

ö

ç

è

æ

-

=

=

F

=

f


· Weekly neutron capture gamma-ray dose equivalent at the door, Hcg:

[image: image145.wmf]week

Sv

Wh

H

cg

m

f

209

10

*

09

.

2

10

*

58

.

2

*

810

4

7

=

=

=

=

-

-


· Neutron dose equivalent at the door, Hn,D:
S0 = inner maze cross sectional area = 15´10´´ x 12.42´ = 18.31 m2
S1 = cross sectional area along maze = 15´10´´ x 7.17´ = 10.58 m2
The Tenth Value Distance for the Maze will be determined by:

[image: image146.wmf]m

S

TVD

7

.

6

58

.

10

06

.

2

06

.

2

1

=

=

=


The neutron dose equivalent will be calculated by the method of Wu & McGinley (2003):


[image: image147.wmf](

)

ú

û

ù

ê

ë

é

+

=

ú

û

ù

ê

ë

é

+

F

=

-

-

-

-

-

-

7

.

6

61

.

6

9

.

1

61

.

6

9

15

9

.

1

1

0

15

,

10

10

*

64

.

1

58

.

10

31

.

18

10

*

27

.

6

*

10

*

4

.

2

10

10

*

64

.

1

10

*

4

.

2

2

2

TVD

d

d

A

D

n

S

S

H



[image: image148.wmf]Gy

Sv

Gy

Sv

H

D

n

m

05

.

2

10

*

05

.

2

6

,

=

=

-


The weekly dose equivalent Hn at the door due to neutrons is W* Hn:

[image: image149.wmf]week

Sv

WH

H

D

n

n

m

1660

05

.

2

*

810

,

=

=

=


· Total Dose Equivalent at the Maze Door Hw:

[image: image150.wmf]week

Sv

H

H

H

H

n

cg

TOT

w

m

2258

1660

209

389

=

+

+

=

+

+

=


· Shielding Barrier for Maze Door:
The maze entrance is located in a controlled area and the shielding design goal is 0.1mSv/week.  The TVL for scatter and leakage photons, 
[image: image151.wmf]TOT

H

, varies between 3-6 mm of Pb depending on maze length.  The TVL for neutron capture gamma rays, 
[image: image152.wmf]cg

H

, can be as much as 61 mm of Pb (NCRP, 1984).  We will assume the shielding for the neutron capture rays will suffice for the scattered and leakage radiation.  We will set 30% of our allowable dose to be from neutrons, and 70% from the capture gammas. We will then have:

[image: image153.wmf](

)

73

.

1

30

1660

log

=

÷

ø

ö

ç

è

æ

=

n

H

n


And

[image: image154.wmf](

)

475

.

0

70

209

log

,

=

÷

ø

ö

ç

è

æ

=

g

c

H

n


The TVL for Borated Polyethylene is 45mm.  We then get:

[image: image155.wmf](

)

BPE

mm

H

t

n

"

1

.

3

8

.

78

45

*

73

.

1

=

=

=


The BPE is on the inside of the door, followed by the lead.  3´´ of BPE allows about 77% transmission of capture gamma rays of the equivalence of 7.1mm of Pb.  We find we need:


[image: image156.wmf](

)

Pb

mmPb

mmPb

H

t

g

c

"

86

.

0

88

.

21

1

.

7

98

.

28

98

.

28

61

*

475

.

0

,

=

=

-

=>

=

=


This calculation neglects the steel casing on the door, which is ¼´´ on each side for a total of ½´´ of steel.  The TVL of steel is 0.108m (McGinley) or 4.25´´.  The amount of radiation that is attenuated by this is:

[image: image157.wmf]23.7%

10

1

10

1

%

25

.

4

50

.

.

=

÷

÷

ø

ö

ç

ç

è

æ

-

=

÷

÷

ø

ö

ç

ç

è

æ

-

=

-

-

TVL

t

attenuated


Our allowed dose through the door was 0.1mSv, 70% of which are capture gammas that are also attenuated by the steel.  Therefore, we have overshielded the door by:

[image: image158.wmf]Sv

mSv

mSv

H

m

6

.

16

0166

.

237

.

*

70

.

0

*

1

.

0

=

=

=


This dose is the extra allowable dose that will added to point E.  

t










34


_1248776979.unknown

_1249195595.unknown

_1249198189.unknown

_1249201040.unknown

_1249209058.unknown

_1249209570.unknown

_1249209590.unknown

_1249209595.unknown

_1249209074.unknown

_1249201402.unknown

_1249202054.unknown

_1249208775.unknown

_1249207113.unknown

_1249202049.unknown

_1249201416.unknown

_1249201353.unknown

_1249201380.unknown

_1249201053.unknown

_1249199063.unknown

_1249200474.unknown

_1249200589.unknown

_1249200626.unknown

_1249201005.unknown

_1249200647.unknown

_1249200612.unknown

_1249200507.unknown

_1249200576.unknown

_1249200499.unknown

_1249200367.unknown

_1249200413.unknown

_1249200314.unknown

_1249198307.unknown

_1249198354.unknown

_1249198408.unknown

_1249198548.unknown

_1249198330.unknown

_1249198294.unknown

_1249197903.unknown

_1249198121.unknown

_1249198185.unknown

_1249198010.unknown

_1249197755.unknown

_1249197868.unknown

_1249197633.unknown

_1248854026.unknown

_1248854634.unknown

_1248856942.unknown

_1248862818.unknown

_1248862846.unknown

_1248862855.unknown

_1248862834.unknown

_1248862342.unknown

_1248856924.unknown

_1248856936.unknown

_1248856920.unknown

_1248854278.unknown

_1248854426.unknown

_1248854261.unknown

_1248777308.unknown

_1248777655.unknown

_1248780091.unknown

_1248780153.unknown

_1248780175.unknown

_1248781173.unknown

_1248780103.unknown

_1248777782.unknown

_1248777816.unknown

_1248777749.unknown

_1248777400.unknown

_1248777643.unknown

_1248777364.unknown

_1248777276.unknown

_1248777289.unknown

_1248777070.unknown

_1248077401.unknown

_1248158891.unknown

_1248179213.unknown

_1248776387.unknown

_1248776553.unknown

_1248776765.unknown

_1248776410.unknown

_1248242687.unknown

_1248242787.unknown

_1248179245.unknown

_1248182014.unknown

_1248179227.unknown

_1248160709.unknown

_1248178618.unknown

_1248179003.unknown

_1248179006.unknown

_1248178660.unknown

_1248161015.unknown

_1248163399.unknown

_1248174462.unknown

_1248177313.unknown

_1248172237.unknown

_1248160718.unknown

_1248160317.unknown

_1248160325.unknown

_1248159557.unknown

_1248160274.unknown

_1248093242.unknown

_1248094192.unknown

_1248095995.unknown

_1248158774.unknown

_1248158827.unknown

_1248157786.unknown

_1248095918.unknown

_1248093530.unknown

_1248092404.unknown

_1248092862.unknown

_1248092913.unknown

_1248092984.unknown

_1248092832.unknown

_1248088979.unknown

_1248090670.unknown

_1248086260.unknown

_1248009910.unknown

_1248074669.unknown

_1248077293.unknown

_1248077383.unknown

_1248077253.unknown

_1248010006.unknown

_1248071096.unknown

_1248009933.unknown

_1248009919.unknown

_1248008011.unknown

_1248009600.unknown

_1248009614.unknown

_1248009715.unknown

_1246876974.unknown

_1248007772.unknown

_1248007892.unknown

_1248007725.unknown

_1246869344.unknown

_1246871069.unknown

